## Slide 1: A Vision-Based Motion-Detection Sytsem Combining MediaPipe with Optical Flow for Long-Term Coma Patient Monitoring
Author: Verena Grall
Supervisor: Dr. D. Pamela

## Slide 2: Agenda
Background & Motivation
Objectives & Requirements
Methodology
System Architecture & Implementation
Validation, Results & Discussion
Deployments, Risks & Conclusion

## Slide 3: Background & Motivation
Clinical need
  detect rare, brief hand movements after long none-movement periods
  Movement indicates, that the comatose patient is awake
Typical computer vision pipelines assume frequent/continuous motion
  Foreground/action methods assume abundant motion
Design goal: precision in a small, semantically relevant ROI

## Slide 4: Objectives & Non-Goals
Objectives
Detect hand-localized motion in real-time
Unified trigger: alert + CSV + clip (pre-roll)
Debounce onset; cooldown determination; re-arming
Portable, CPU-only; deterministic & auditable
Non-Goals
Clinical decision-making or diagnosis
Full-body/multi-patient tracking
Cloud services
Training new models in-scope

## Slide 5: Key Requierments
Functional
Single-hand ROI
Dense optical flow in ROI
Debounce triggers; pre-roll clips
Cooldown; re-arming; pause/grace
Clean shutdown
Non-Functional
Cross-platform codecs and alerts
Timestamped filenames; CSV audit log

## Slide 6: Method Overview
Localize hand ROI with MediaPipe Hands
Frame-level decision by area fraction
Temporal debounce (>=3 frames) before trigger
Compute Farnebäck dense optical flow in ROI
Threshold magnitude; morphology to denoise

## Slide 7: ROI Localization
Blue rectangle marks the hand-centric ROI (padding = 10 px).

## Slide 8: Motion Blobs with ROI
Green rectangles mark thresholded, morphologically cleaned motion inside the ROI

## Slide 9: When motion is detected …
Trigger recording:
  Open VideoWriter
  Add pre-motion frames
  record current frames until no motion + cooldown elapses
Trigger alert
  Winsound on windows
Trigger CSV logging
  Write new row in CSV with timestamp, filename and event
Re-arm after quiet period

## Slide 10: User Interface
Output frames are shown on screen
Stop with ESC-, ‚q‘-key, or closing the window
Pause/resume with ‚p‘-key
  During resume grace, no motion is detected

## Slide 11: Implementation Details
Python 3.x – OpenCV – MediaPipe – NumPy
Cross-platform codecs: XVID/AVI (windows), MP4V/MP4 (others)
Alert sink: winsound (windows) or WAV fallback
CSV logging with header auto-creation

## Slide 12: Validation Plan (Tests)
V1: File-based determinism (idle baseline; scripted bursts)
  Expect 0 events in idle baseline
V2: Live-loop sanity (re-arming; pause/grace behavior)
  1 alert + 1 CSV + 1 clip per burst; pre-roll & tail check
V3: Edge & stress (ROI loss; planned: flicker, sweeps)
  Metrics: trigger correctness, latency, integrity, stability

## Slide 13: Results: Scripted-Burst Integrity
Clip | fps_nom | fps_real | rec_s | pre-roll (exp) | tail_s (min) | pre-roll_ok | tail_ok | PASS
movement_...38-04.avi | 30.01 | 28.48 | 6.81 | 60 (60) | 5.06 (5.00) | True | True | True
movement_...38-16.avi | 30.01 | 27.92 | 7.27 | 60 (60) | 5.05 (5.00) | True | True | True
movement_...38-29.avi | 30.01 | 28.26 | 8.74 | 60 (60) | 5.07 (5.00) | True | True | True
movement_...38-43.avi | 30.01 | 28.47 | 6.74 | 60 (60) | 5.04 (5.00) | True | True | True
movement_...38-56.avi | 30.01 | n/a | n/a | 60 (60) | 3.63 (5.00) | True | False | False

## Slide 14: Results
Trigger correctness: 100% (5/5 events)
  FAIL explained by video ending before cooldown
Pre-roll frames: 60/60 in all valid events
  Lifecycle aligns alert/CSV/clip to same onset
Tail (time): mean 5.055 s, SD 0.011 s (4/5 valid)
  Design meets integrity targets under constraints
Realized FPS: ~28–29 vs nominal 30; expected on CPU

## Slide 15: Risks & Mitigations
Key Risks
  Missed subtle motion (false negatives)
  Lighting flicker inside ROI (false positives)
  ROI loss/occlusion
  Storage exhaustion; muted alerts
Mitigations
  Tune threshold/kernel/FPS
  Increase arm frames; stabilize lighting
  Static-image mode; padding; last-seen ROI
  Prune by age/size; WAV/OS notify fallback

## Slide 16: Conclusion
ROI-restricted dense flow + debounce/cooldown works for sparse events
Unified trigger yields synchronized alert/log/clip with pre-roll
Engineering validation supports lifecycle correctness
Future: multi-ROI (hands, feet, eyes), adaptive sensitivity, alerting through e-mail/WhatsApp/SMS

## Slide 17: Thank You
Verena Grall